## Slide 1
SoI — MVP Stage
Submission Deck Template
Team Details :
Team: The Spark | Team-ID: P-2024-28-EI-159
Lead: Sibithra V | Members: Sakthivel V, Arun R, Viviksha K, Shri krishnamoorthy S
Mentor: Sriram R | Lab Code: EI
MVP Outcomes:
The MVP successfully detected microplastics in water in real time.
The system achieved fast response with acceptable accuracy during user testing.
Some instability was observed due to ambient light and alignment issues.
Further refinement is required to improve accuracy and reliability before deployment.
MVP Stage — Submission Deck
1/23

## Slide 2
SoI Innovation Pipeline
Problem & End-User Reconfirmation
Problem & user
Target End-User : Water treatment plant operators and municipal water quality inspectors.
Pain Point : Current microplastic detection requires lab testing, expensive equipment, and long turnaround time (hours to days).
Context / Environment : On-site water testing benches in municipal plants; indoor testing environment.
Baseline (Current State Metric) : Lab-based testing takes 24–72 hours and costs significantly per sample.
Why Urgent : Microplastic contamination is widespread; delay in detection prevents immediate corrective action
Value proposition
For water quality inspectors who need fast contamination checks, our device provides on-site detection within 2 seconds so they can take immediate action.
Success metric anchor
Time to Detection (seconds) If we don’t reduce detection time drastically, the product is useless.
MVP Stage — Submission Deck
5/23

## Slide 3
SoI Innovation Pipeline
Prototype Recap
What worked in prototype
End-to-end hardware integration achieved
Detection accuracy ~60%
Response time ≤ 2 seconds
Constraints Discovered :
Signal drift over time
Ambient light interference
Alignment sensitivity
Top gaps blocking MVP
Need better enclosure
Need calibration mechanism
Need real-user validation
What changes at MVP
Real usage by lab users
Structured feedback collection
Metric composition
At least one iteration after feedback
MVP Stage — Submission Deck
6/23

## Slide 4
SoI Innovation Pipeline
MVP Hypothesis (Mandatory)
Write your hypothesis
We believe that water quality technicians will reduce microplastic detection time from hours to under 2 seconds if they use our portable optical detection MVP in a lab or field setting.
Success criteria (minimum 2 metrics)
Metric
Target
How measured
Detection Accuracy
30 test runs (clean vs contaminated)
Detection Time
≤2 sec
Stopwatch measurement
False Alerts
≤5%
20 room-light tests
Tier target
Tier 2 : 5–10 users + ≥2 KPIs + 1 iteration
MVP Stage — Submission Deck
8/23
≥95%

## Slide 5
SoI Innovation Pipeline
MVP Scope
MVP user journey
MVP User Goal
Quickly check if a water sample contains microplastics.
User Journey
Power ON device
Insert water sample
View contamination status on LCD
Output
“Water clear” or “Particles Found” within 2 seconds.
In scope
Optical sensing
Signal processing
LCD contamination alert
Basic enclosure
Out of scope
Plastic type classification
Cloud Dashboard
Mobile app
Mass deployment
MVP Stage — Submission Deck
9/23

## Slide 6
SoI Innovation Pipeline
Deployment Plan
Where + who
Deployment Sites
College environmental lab + 1 external water testing lab
End-User Group
Lab assistants / technicians
Planned User Count
6 users
Duration
5 days
Access & onboarding
1-page instruction sheet
Live demonstration
No login required
Safety / privacy / ethics
MVP Stage — Submission Deck
10/23
Written user consent
No personal data collected
Insulated wiring
Secure enclosure

## Slide 7
SoI Innovation Pipeline
REAL vs SIMULATED (Mandatory)
Component / step
Real
Simulated
End-user interaction
✔
-
Real lab users
Data capture
✔
-
Arduino logs
Core logic
✔
-
Threshold algorithm
Integration
✔
-
Full hardware
Notifications
✔
-
LCD display
Note: Simulation is allowed only if clearly marked. MVP still requires real end-users and real feedback.
MVP Stage — Submission Deck
11/23

## Slide 8
SoI Innovation Pipeline
User Validation Plan
Test protocol
Users Recruited
6 lab users
Tasks
Test clean water
Test plastic-mixed water
Repeat under room lighting
Time per Session
5 minutes
Feedback questions (minimum)
Quantitative quick measures
MVP Stage — Submission Deck
13/23
Was the result clear?
Was setup easy?
Any confusion?
Would you use this regularly?
Task success rate
Detection time
False alert rate

## Slide 9
SoI Innovation Pipeline
Feedback Summary
What users said
Worked Well
Very fast results
Simple display
Easy to operate
Did Not Work
Sensitive to bright lighting
Needs protective casing
Confusing
Threshold calibration process
Top quotes
“Very fast compared to lab methods.” — User 1 “Needs enclosure for stability.” — User 2
Adoption signals
Would use again: 5/6 Recommend score: 8/10 average
MVP Stage — Submission Deck
14/23

## Slide 10
SoI Innovation Pipeline
MVP Metrics (Target vs Actual)
Metric
Target
Actual
Status
Evidence (slide/link/log)
Detection Accuracy
95%
85%
Pass
Test log – 20 sample runs
Detection Time
≤2 sec
1.8 sec
Partial
Demo video timing
False Alerts
Pass
Room light test log
Reliability
1 hr stable
1 hr stable
Partial
Continuous run video
MVP Stage — Submission Deck
15/23
7%
≤5%

## Slide 11
SoI Innovation Pipeline
Iteration Log
Version
Change made
Why
Impact
Prototype
Basic setup
Initial testing
MVP v1
Enclosure added
Reduce light noise
90% accuracy
MVP v2
Signal averaging improved
Reduce false alerts
MVP Stage — Submission Deck
17/23
85% accuracy
Faster response

## Slide 12
SoI Innovation Pipeline
Pain-Point Proof
Original pain point
MVP evidence
Status
Slow and costly microplastic detection.
Detection time reduced from hours → 2 seconds.
Partial
Conclusion: Based on MVP evidence, our solution significantly reduces detection time but needs accuracy improvement before large-scale deployment.
MVP Stage — Submission Deck
18/23

## Slide 13
SoI Innovation Pipeline
Evidence Bundle (What to capture)
Screenshots / photos
MVP in use (with consent)
Key screens / flows
Hardware deployed / setup
Error cases (if any)
Logs / data / forms
Usage logs / session counts
Key Performance Indicator (KPI) measurements table
Feedback form export (anonymized)
Consent record (sample)
Iteration notes
MVP Stage — Submission Deck
19/23

## Slide 14
SoI Innovation Pipeline
Demo Video (MP4 ≤ 2 minutes) + Artifacts
2-minute demo script
0:00 – Problem: Microplastics need lab detection. 0:20 – Show device setup 0:40 – Test clean water 1:00 – Add plastic → show detection 1:30 – Show metric comparison 1:45 – Mention feedback + iteration
Links
Repo link: https://github.com/Sibithra/Detection-of-Microplastics-in-Water-
Demo video link : https://www.kapwing.com/videos/69a4f7c34c95423eb3e0bf77
MVP Stage — Submission Deck
21/23

## Slide 15
SoI Innovation Pipeline
Risks, Next Steps & Direction
Top risks & mitigations
Ambient light interference → optical chamber
Signal drift → auto-calibration
Limited accuracy → advanced filtering
Decision
REFINE (partial)
Next steps
Improve enclosure
Increase test sample size
Improve threshold algorithm
Target ≥95% accuracy
MVP Stage — Submission Deck
22/23

## Slide 16
SoI Innovation Pipeline
References
Add your references here
[1] NIH Research Matters, “Plastic particles in bottled water,” National Institutes of Health, 2024. [Online]. Available: https://www.nih.gov/news-events/nih-research-matters/plastic-particles-bottled-water. [Accessed: Jan. 20, 2026].
[2] A. Shrivastava et al., “Microplastics in drinking water: Detection and challenges,” Sensors, vol. 4, no. 2, 2024. [Online]. Available: https://www.mdpi.com/2673-8929/4/2/24. [Accessed: Jan. 20, 2026].
[3] Deutsche Welle, “Bottled water not safe from microplastic contamination,” 2018. [Online]. Available: https://www.dw.com/en/bottled-water-not-safe-from-microplastic-contamination/a-42936246. [Accessed: Jan. 20, 2026].
[4] S. Mintenig et al., “Low numbers of microplastics detected in drinking water,” Water Research, vol. 142, pp. 488–495, 2019. [Online]. Available: https://www.sciencedirect.com/science/article/pii/S0043135419301794. [Accessed: Jan. 20, 2026].
[5] PubMed, “Microplastic exposure via drinking water,” PubMed ID: 38537794, 2024. [Online]. Available: https://pubmed.ncbi.nlm.nih.gov/38537794/. [Accessed: Jan. 20, 2026].
MVP Stage — Submission Deck
23/23